Redesign of OrthoLPPDUNS External Fixation for Bone Reconstruction Using the Function Analysis System Technique (FAST) Method

Lobes Herdiman1a, Ilham Priadythama1b, Susy Susmartini1c, Yusuf Priyandari1d, 
I Nyoman Suci Anindya Murdiyantara2e, and Durkes Herlina Apriani1f

Abstract. Taylor Spatial Frame (TSF) better known as an external fixation product from several types of fixation devices in the orthopedic field. The Laboratory of Product Planning and Design has manufactured TSF or called OrthoLPPDUNS external fixations. The OrthoLPPDUNS design found that frame instability forming ring-strut angles caused by strut screw rotations was felt rough due to backlash. The purpose is how to redesign an OrthoLPPDUNS with a Function Analysis System Technique (FAST). FAST method is applied at each stage of starting from the information stage to the interpretation of the results. The OrthoLPPDUNS design with the joint system uses custom components that cause weakness in frame stability. The use of standard components in joint systems can provide solutions to external fixations, although there are still weaknesses in the ring-strut angle. The OrthoLPPDUNS redesign has a better success rate in terms of stability parameters, allowing more precise bone reconstruction.

Keywords: OrthoLPPDUNS; FAST; frame fixation; stability.

# Introduction

External fixation techniques were popularized in the mid-20th century when Hoffman introduced a tool that used bars Steinman and half-pins to stabilize long fracture bones (Drijber et al., 1990). External fixation is a metal frame that is designed to be able to hold and stabilize bones in place (Dubowy and Nichols, 2018). The use of external fixation to reconstruct bone limbs due to fractures or abnormalities in the bone and has an adjustable bar (called strut) is rotated slowly in reconstructing the bone. In other words, external fixation is a method for treating bone or joint injuries and for correcting bone deformities by attaching external devices to the bones that aim to stabilize injured limbs (Solomin, 2008). In addition, it allows for the manipulation of limb segments to do planned length and alignment.

In 1989 Behrens explained that three basic concepts governing the use of external frames safely and effectively for bone trauma namely K-wire and half-pin are vital mechanical structures and must be protected from damage, allow access to the area of injury, and must meet mechanical demands patient. In general, there are three types of external fixations that are widely used in bone reconstruction (Solomin, 2012). Ilizarov's external fixation is a circular fixation consisting of a metal ring or carbon fiber that is connected by metal bars and clickers (Fig. 1a). The Ilizarov fixation is connected to the bone with thin wire and thick half-pins that are used to hold the bone in place. A monolateral external fixation is a fixation bar that consists of metal bars connected to bones with thick pins (Fig. 1b). This device is used to extend or hold the bone in place after it is straightened. Use screws to hold the bar in place when the bone extends over the bar. Taylor's Spatial Frame (TSF) is a circular fixation consisting of two metal rings and six telescopic struts at universal joints, creating a hexapod device system (Fig. 1c). TSF is a circular external fixation system, allowing correction of skeletal deformity from the simplest to the most complex using the same frame. This frame has two full circles or two per thirds of the rings connected by six struts. Use thin wire and thick half-pins to connect to the bone.

In 1998 H.S. and J.C. Taylor first created the Taylor Spatial Frame (TSF, Smith & Nephew, Memphis, TN, USA) and applied this method to orthopedic (Taylor H.S. and J.C., 2000; Tan et al., 2014). This frame consists of a ring and a strut (Bone Fixator, 2018), a ring made of aluminum, titanium, or carbon fiber, available in various ring sizes determined by internal diameters, 80, 105, 130, 155, 180, 205, 230, 255 and 300 mm. Strut is available in four long groups namely extra short (75-96 mm), short (90-125 mm), medium (116-178 mm) and long (169-283 mm). The attachment of the frame to the bone is equipped with wire from stainless steel or titanium alloy with a diameter of 1.6 or 1.8 mm (called K-wire) that penetrates the skin and enters the bone. The use of this frame is paired with a half-pin (called a pin) on the ring that penetrates the skin and enters the bone using stack rancho cubes, available in diameters of 4, 5 and 6 mm. The two ring construction (top for proximal and bottom for distal) can simulate single-level deformity; a three-ring construction with six struts in each ring pair can simulate two-level deformity (Paley, 2011). Modification of strut lengths by changing strut sizes in frame configurations allows changes in the orientation of one ring, (Keshet and Eidelman, 2017) leading to simultaneous correction of each deformity (differences in leg length, angulations, translation, and rotation). The position of the computer-generated ring is based on measurements taken from X-rays and clinical examination. After all deformity parameters in the software, a daily set schedule can be proposed and printed, a strut is rotated every day to straighten or extend the bone.

Construction of a six strut angle on the TSF with the axial weight of each strut without applying bending forces to the tilt of the strut attached to the ring When we look at the attachment points strut on the ring, they are shaped like a triangle, not a circle. Paley, 2011 explained that all structures in this construction, strut configurations that form side triangles, have the same shape as the diamond crystal structure (octahedron). This construction is a very strong construction. Manual mechanical accuracy for adjusting the position of the strut in the correction of the six-axis deformity has been measured up to 0.7 degrees and 2 mm. The TSF fixation is able to correct all aspects of the six-axis deformity simultaneously (Keshet and Eidelman, 2017). Components of deformity and divide it into angulations, rotation, translation, and shortening or lengthening (Heidari et al., 2013). Angulations and rotation are angular deformities, measured in degrees. Translation and length are displacements, measured in units of distance (e.g., millimeters, inches,). According to Paley (2002) explains that the six parameters of deformity needed to define clinical deformity include anterior-posterior plane angulations, lateral plane angulations, axial plane angulations, and anterior-posterior plane translational, lateral plane translational and axial plane translational. Correction of deformity between two bone segments can be characterized by three projected angles (rotation) and three projected displacements (translational). Therefore, six parameters of deformity are needed to decide the place of bone deformity in a patient.

In 2015 researchers from Industrial Engineering the Laboratory of Product Planning and Design at Universitas Sebelas Maret Solo and an orthopedic doctor succeeded in making a TSF replica product called OrthoLPPDUNS external fixation or OrthoLPPDUNS. The making of this product is to assist orthopedic doctors in correcting tibia bones in a patient with Blount's disease or known as another tibia vara at Rumah Sakit Unggulan Pendidikan (RSUP) Kabupaten Klaten Jawa Tengah, Indonesia. These patients need osteotomy surgery to correct Blount's disease, the occurrence of this deformity can occur in infancy, childhood, and adolescence (Putra, 2018; Sabharwal, 2016). Osteotomy surgery in these patients aims to improve deformity, equalize the length of both legs, and improve symptoms. Acute corrective osteotomy in patients has been complicated by peroneal nerve palsy, compartment syndrome, residual deformity (Eralp et al., 2016), long limb imbalance (Henderson et al., 1992), delayed unification of bone, or events with emergence fixation device failures (Birringer et al., 2016). The results of a review of several journal literatures, that the use of circular frames has shown the ability to gradually correct angulations, translation, rotation, and lengthening or shortening, thereby minimize the incidence of postoperative complications.

Before an osteotomy surgery is performed on the patient, the OrthoLPPUNUNS device is first assessed on the joint system at the ring-strut angle. The results of the study on the design of the OrthoLPPDUNS are made by custom components through the manufacturing process. After a number of motion simulation tests using the Fusion 360 software, it was found that the rotational motion with displacement values for strut screw components appeared to be rougher or not smooth. The reason is that the threads on the strut components are worn out. The found of this weakness results in unstable frames at a certain distance, seen from the mass of the mass with the use of an inner ring diameter of 155 mm and a combination of six struts with a medium length of 
150 mm. The configuration of the frame at the ring-strut angle shows this instability due to backlash. The appearance of backlash on the joint system between the ring-strut angels due to the external fixation design requires a clearance plane between components. This design results from considering the TSF design patent no. US6,030,386 (Taylor H.S. and J.C., 2000) which explains that giving clearance is required for the connection between components on the edge shoulder bolt and the edge of the universal joint, applied to every component of the universal joint.

The provision of clearance between components aims to enable the frame configuration to move in a three-axis system (Shah et al., 2012), whereas the universal joint component is a two-axis joint system (Wijaya et al., 2019). Giving clearance allows strut mounted on the ring under tight conditions, but can still be rotated by the patient. However, giving clearance out of control will cause a backlash and make the strut twist process heavy and the correct size specified in the previous around schedule inaccurate. If this condition is not a concern, it will cause undesired bone shifts during bone reconstruction in the patient. Although the process of bone reconstruction in a patient is successful, the results of correction remain inaccurate. Several studies have evaluated the parameters that affect the stability of the circular fixation. These factors such as ring size and material; wire type, diameter and voltage; the number, size, and location of half-pins have been shown to have the same effect on the stability of TSF fixation. However, studies of frame stability related to ring-strut angles have not been reported in the literature. We redesigned this device, gave to attention to changes instability in the joint system between the ring-strut angles due to the use of alternative components on the external fixation using the Function Analysis System Technique (FAST) to overcome backlash and struts screw that wear out quickly.

# RESEARCH METHOD

Techniques in limb reconstruction continue to show significant progress, largely due to the use of external fixation devices. The emergence of backlash and strut screw problems that wear out quickly requires the redesign stage in OrthoLPPUNUNS with the discussion covering the information stage, the functional stage, the evaluation stage, the redesign stage and the design interpretation stage (Figure 2).

Information stage

The information phase was carried out by means of interviews through discussions with orthopedic surgeons who had recognized and operated the external fixation limb of RSUP Dr. Soeradji Tortonegoro Klaten and the external fixation product design team at the Industrial Engineering from Laboratory of Product Planning and Design at Universitas Sebelas Maret Solo. Look up information from type of TSF external fixation functions, and describe the workings of the device. The results of this discussion will be obtained from an overview and notes from the OrthoLPPDUNS design. This information is then stated in the form of a description of the need to ease the preparation of the steps when redesigning the external fixation.

Function stage

The description of the external fixation mechanism using FAST method aims to map and sort the functions of each frame component during operation. This method also aims to visualize the relationship between all functions in the frame, making it easier to check how the frameworks. Compilation of functions with two questions between how and why, it’s namely how the external fixation works and why a bone correction is performed. In the first stages, the FAST method is described in the flow diagram FAST is made in general to decide the work function of each component as to why a bone correction is done.

Evaluation stage

At this stage, the design of OrthoLPPDUNS is examined in more detail. Through evaluation of the flow of FAST diagrams on each frame component function, so that the component functions can be traced clearly to weaknesses in the design. Determination of technical requirements, as recommendations for what it’s should be done in the redesign of the external fixation which is a consideration in bone reconstruction.

Redesign stage

It’s having known the weaknesses of the use of components in the OrthoLPPUNUNS design. Then the external fixation design is compared in the use of components and provides recommendations using other components. The external fixation design is explained in the form of three-dimensional images using Autodesk Inventor 2016. The FAST diagram flow again explains the external fixation redesign using other components as a replacement.

Design the interpretation stage

This stage is conducted on a simulation test on OrthoLPPDUNS design between custom components and standards. Compares the characteristics of the two frames based on physical dimensions and stability parameters including the value of the safety factor, displacement and the range of the ring-strut angle. Constraints and further study of the external fixation design to get the best ring-strut angle.

# RESULT AND DISCUSSION

The information phase aims to decide the user needs of doctors and patients from the OrthoLPPDUNS design. This study involved an orthopedic surgeon and two designers from an external fixation device at the Laboratory of Product Planning and Design. OrthoLPPDUNS products are custom-made with all components through a manufacturing process with manual machining. The results of discussions with an orthopedic physician involved in this study illustrate in his explanation that the OrthoLPPDUNS design focuses on the ability to rotate the screw strut. This indication is feared that the frame will experience undesired movements after the strut head rotations by the patient. The frame in the OrthoLPPDUNS design will experience mechanical instability originating from the achievement of the ring-struts angle. Henderson et al., 2008 recommend that long struts be selected when installing TSF so that the angle of the ring-strut can reach 30 degrees, avoided less. Clinical relevance is related to the use of TSF in atypical confirmation, especially in the pediatric population who need correction of the deformity. After obtaining a response from an orthopedic doctor and information from several journal articles, a thorough review of this design was then carried out with the external fixation design team. The results of this study note that there are still indications of the emergence of undesired movements originating from the coarse motion of the strut screw components (Figure 3).

In addition, the results of a motion simulation test using the Fusion 360 software show that there are several strut components in the strut screw section that look worn; this condition affects the movement of the frame. The next step is how to redesign OrthoLPPDUNS with the rotating motion of the strut screw being smooth and the threaded component does not wear out quickly. The two stability parameters in this frame are needed for immediate repair to produce a correct bone correction.

Analysis of the function stage of the external fixation component using the FAST method is related to when operating the frame, followed by arranging these functions into the FAST flow diagram. The first step in this analysis is to describe the flow of the FAST diagram in an external fixation in general (Figure 4).

The operation of the frame external fixation begins with how to lengthen and shorten the struts. Where turning the struts cap to the right (+) will increase the struts screw and push down the struts tube, then the struts will be elongated, otherwise, the struts cap turning to the left (-) will decrease the struts screw and pull the struts tube, the struts will shorten. Changing the length of the struts can change the position of the joint; the angle of the joint can be changed. Changing the joint angle changes the position of the shoulder bolt attached to the ring so that the place of the ring changes. Changing the position of the ring will automatically change the place of the 
K-wire fixation bolt mounted on the ring, which will affect the position of the K-wire. Changes in bone place due to K-wire attached to the bone. The operation of this frame is done periodically until a bone correction can be achieved according to the planned schedule. FAST diagram flow can be seen that the stability factor in the frame component between the ring-strut for external redesign fixation is an important component.

In the evaluation phase by depicting the flow of the FAST diagram on the external frame fixation then tracing the root of the problem that occurs when operating the frame relating to frame stability (Figure 5).

OrthoLPPDUNS design, the universal joint is a connecting component between ring-strut angles which is placed in the strut section. The results of the motion simulation which states that the rotational motion of the strut becomes coarse can be traced from the flow diagram below (Figure 6).

The OrthoLPPDUNS design frame uses custom universal joint components. Vesali et al. (2012) explain that the universal joint component consists of two hinges connected by a crossbar, allowing the bar to turn in any direction (Figure 7). Adjusting the position of the frame on the strut will change the angle of the ring-strut; this change is influenced by the direction of strut rotation. At the time of a strut tube twist arises as if the rotation is not smooth.

This condition occurs due to backlash on the connection component. Backlash according to Solon (2010) explains the required distance or space or the mess between two or more components. Backlash at the connection between the ring-strut is required for clearance in the center block hole between the small pin and large pin. In addition, clearance is also needed at the end of the universal joint connected to the ring. Clearance space requirements on components a small size to support the mechanism of motion of the mechanism on the external fixation (Karidis, 2012). However, clearance with excessive size causes arousal to occur in the frame especially strut components that are related to the universal joint between the center block and pin, as well as the joint and ring. According to Seherr-Thoss et al., 2006 explains that excessive clearance will reduce the accuracy of the dimensions. The universal joint in the OrthoLPPDUNS design is custom with a single component, where the dimensions of the yoke and connections in each component are not adjusted according to standard products available in the market (Figure 8).

The purpose of a custom universal joint is that the angle of the ring-strut formed is greater so that tolerance between components is not required even though as a result the distance between the components becomes greater. So, it is increasingly clear that the problem of strut twist movements of frames that are not smooth due to backlash from the use of custom universal joints.

Position the displacement in the frame by lowering and raising the strut screw so that the joint angle of the ring-strut can be changed. When strut tube playback is not smooth within a certain time the strut screw will wear faster (Figure 9).

The screw rod design on the strut component is a custom screw from another manufacturer's external fixation dimensions. The strut screw (threaded rod) design is not according to product standards available in the market but is a custom product. The screw pitch in the OrthoLPPDUNS design is 1 mm with a major diameter of 5 mm, designed so as not to affect the twist system. The design of the threaded rod is not yet seen as a weakness, after the simulation test by evaluating the displacement value shows the greater value. This test involves turning parameters of the manual threading process at a certain time with several iterations in the simulation without tool changes, that the diameter of the rod is reduced and the results are inconsistent. Even though inaccurate dimensions of threaded rod size will cause continuous friction on the strut cap component, the threaded strut becomes easily worn. According to the standard metric screw thread (ISO 724, 1993) explains that the pitch size of the threaded rod is 1 mm with a major diameter of 6 mm. Allowing this error potential to be greater when of production let alone lack of supervision at every step of the production process (Evans, 2014). So, it is increasingly clear that the problem of the strut screw is easily worn due to the use of custom screw components.

Both of these problems can be identified and are considered as recommendations for OrthoLPPDUNS redesign of frame stability parameters, namely joint component specifications with low or even zero backlashes, use of screws that do not wear out quickly, and a strong ring construction design.

Redesign stage for the need for an external frame fixation, where strut movements must be smooth and avoid excessive backlash events. After was discussion with orthopedic doctors and practitioners in the field of designing an external fixation. The concept of repair is done by replacing custom universal joint components using standard components with ball joint components, where the selected joint is spherical bearing. The selection of spherical bearings is based on the SKF catalog because it is able to make rotational movements at a wider angle, suitable to be a substitute for a custom universal joint. The spherical plain bearing is a standard mechanical component that allows movement in various directions and is able to align the bearing itself. This bearing consists of two components, an inner ring that has an outer surface such as a ball and an outer ring that has a concave-shaped inner surface that adjusts the inner ring (SKF, 2011). The advantages of spherical plain bearings as an external joint fixation are able to accommodate misalignments, high reliability because there is almost no pressure from the edge and the load exceeds the limit, accommodate the deformation of the surrounding components, accommodate wide manufacturing tolerances, able to carry high capacity loads, suitable used for slow and precise rotating movements, long economic life (INA, 2008). The bearing used is radial spherical plain bearing with GE 8E type steel material (Figure 10).

Redesign of the strut component is that the screw does not wear out technically, it is stated that the screw has high-quality material. The concept of repairing screws is done by replacing custom threaded rod components with standard threaded rods, the standard metric screw threads used according to ISO 724 (ISO, 1993). A standard threaded rod which has a pitch of 
1 mm with a major diameter of 6 mm. This standard product has advantages, where manufacturing efficiency, reliability and quality are guaranteed (Evans, 2014) and easily obtained at a cheaper price (Nicholas, 2010). Replacement of components with standard products is done in the hope that these components have a guaranteed quality standard size and are not easily worn or damaged. The results of the evaluation stage in the OrthoLPPDUNS design for the joint system can be determined by comparing the use of custom and standard components (Figure 11 and Figure 12).

The ring component is part of the frame is the main component that serves as a 
K-wire buffer made of aluminum 6061A material that is radiopaque (difficult to penetrate X-rays). Frame construction is made in a circle with the aim to produce stability in the frame. The hole in the outer ring is used to install the strut components; the inner ring hole is used to install other components such as K-wire and half pin attached to the bone to support the rigidity of the frame (Smith & Nephew, 2011). The strut component is a part of the frame consisting of a cylinder with a threaded rod in it, where the rod can be raised or lowered on the cylinder by turning the strut cap. The threaded rod used in this fixation a pitch of 1 mm and a major diameter of 5 mm. One turn of the strut head can change the length of the strut by 1 mm. This component is connected to the universal joint component, where the strut at the bottom is modified to be yoke 1. The threaded rod component is made of stainless steel material and the cylinder component and strut head are made of al. 6061A which is radiopaque. The length of the strut can be changed at certain intervals so that the position of one ring can change to another ring (Paley, 2002).

Custom universal joint components are components that are mounted on the end of the strut cylinder and the end of the threaded rod as a component of the ring-strut connector. This component consists of the main cross axle, cross axle pin, and two yokes. Yoke 1 and the cross axle pin are connected in a fix, while yoke 2, the main cross axle and cross axle pins are connected in such a way that they can rotate with each other. Using a joint to support changes in the position of the ring to another ring, this joint system provides the ability to "turn" in any direction, made of al. 6061A which is radiopaque. The shoulder bolt component is a component that connects the yoke 2 part to the custom universal joint against the ring, made of stainless steel with a standard head diameter of 10 mm. The use of shoulder bolt components is done so that the strut can still rotate to the ring even if it is connected tightly enough (Norrish, 2013).

The ring component consists of two rings with a diameter of 155 mm with a thickness of 3.75 mm made of al. 6061A and connected with three countersink screws. This component has a cavity that serves as a bearing housing and a hole with a diameter in the ring to install other components, such as K-wire and half pin. Strut components made of al. 6061A materials, consisting of 
6 strut tubes and 6 strut screws arranged in such a way. Strut screw made of stainless steel material in this design uses a standard M6 threaded rod metric. Bearing components, this design uses 12 spherical plain bearings with a diameter of 8 mm with GE 8E specifications. This bearing is installed in a cavity in the ring. Button head shoulder bolt component that serves to connect the strut with bearings mounted in the ring. The shoulder bolt head is modified into a button shape so that it can be mounted on the bearing without limiting the bearing space.

The FAST flow diagram for the OrthoLPPDUNS design with standard components such as spherical plain bearings and threaded rods can be described again to find out how to operate this frame to achieve bone correction (Figure 13).

The joint system in spherical plain bearings is a ball-type joint that aims to overcome the need for clearance at the joint ends of the ring-strut with three degrees of freedom. The use of standard components on the external fixation has met the needs of smooth strut movements; the need for clearance can be minimized. Spherical plain bearing SKF products with clearance as a benchmark tolerance of connections between components have a connection tolerance of 68 μm and a custom component of 1150 μm. Spherical plain bearings are installed in the ring, where the shoulder bolt is attached to the bearing. The strut is mounted directly on the shoulder bolt and bearing without touching the ring, the installation does not need clearance, the strut can rotate freely. This external fixation uses a standard threaded rod to meet the needs of a threaded strut that is not easy to wear; the risk of wear can be minimized.

The design interpretation stage is simulated through finite element analysis using Fusion 360 software on the OrthoLPPDUNS frame design. This test by giving loading at points on the ring that receive direct loads, where this point is where the installation of K-wire, there are eight points spread on both rings. The magnitude of the load is based on the ability of material al. 6061A with an average adult weight of 700 N with a withdrawal load on the K-wire of 1100 N.

The value of the safety factor in the frame design should be considered when designing the frame. The ability to accept or safety factor is the value of the ability of the structure to design statically in receiving loads, which is measured when the higher the value, the better the acceptance of the load (Ullman, 1986). The simulation in the safety factor test aims at whether the frame is still capable of receiving heavy loads from a predetermined person, whether this frame is able to accept a greater or lesser burden on the allowable strength. Modern engineering that the value’s of the safety factor in design ranges from 1.25 to 4.0. Safety factor simulation results on the standard component frame of 4.21 ul, showing that this external frame fixation has better strength than frames using custom components.

The displacement value indicates the maximum value of the shift that is still allowed in the frame. The displacement value is the benchmark in the design for frame shifting. It’s same as a benchmark for shifting between unwanted bone fragments or Inter-Fragmentory Shear (IFS). Until now the value of IFS is still controversial, some studies suggest this value can have positive side effects. According to Betts and Muller (2014) explained that the majority of studies on the value of displacement in frame fixation stated that IFS can inhibit bone reconstruction. IFS can also delay bone formation in osteotomy between the bones, decrease callus in the periosteal, and stiffness in bone tissue that should be flexible (Augat et al., 2003). Meanwhile, according to Steiner et al. (2014) states that the allowable value of IFS during bone reconstruction is the maximum value of inter-fragmentory movement/IFMmax (movement between bone fragments) of 
1.5 mm. Based on the majority of these studies, an IFS value that exceeds regulatory limits is something that should be avoided, because it can inhibit bone reconstruction. IFS must be calculated to a minimum in order to anticipate obstructed bone reconstruction. The displacement value on the frame within the safe limit is in the range of 0.115 to 0.388 mm (Figure 14).

The use of standard components in the frame external fixation has a small displacement value of 0.115 mm or only 10% of the maximum displacement allowed. These results indicate that this frame is still possible for IFS to occur in bone, but with a very small value for the occurrence of bone reconstruction obstructions. Frames with small displacement values are declared quite safe. In this simulation test, it is explained that green has a small displacement and red has a large displacement. Large displacement frames are on two ear rings, due to the loading point of the K-wire. This loading point is still at a safe level, where the surface area around the loading becomes uniform. The results of the evaluation and redesign recommendations are produced by the OrthoLPPDUNS frame with the use of standard components (Figure 15).

The ring-strut angle reach in the joint system of the OrthoLPPUNUNS uses custom and standard components (Figure 16).

Custom components are capable of forming an angle of 96.08 degrees or a 
0 degree ring-strut angle; the dimensions of the yoke have been modified so that it does not obstruct other yokes when rotating. The maximum angle uses a standard component of 42.08 degrees or a ring-strut angle of 47.92 degrees. Although both of these frame fixations with ring-strut angles formed cannot reach the recommended angle of approaching 30 degrees (Henderson, 2008).

The limitation of the formation of this angle is the impact of mounting the joint in the ring so that the strut can only rotate as broadly as the ring hole. This condition has an impact on the limitations of the free-angle in reconstructing bone later. Although, it problem can be overcome by replacing the longer strut. But strut replacement is also experiencing problems because the frame has been formed with certain strut angles, making the attached shoulder bolt difficult to remove. It should be noted in the next frame study by considering the tolerance of each component in the external frame fixation to avoid backlash again. The limitation of the formation of this angle is the impact of mounting the joint in the ring so that the strut can only rotate as broadly as the ring hole. This condition has an impact on the limitations of the free-angle in reconstructing bone later. Although, it problem can be overcome by replacing the longer strut. But strut replacement is also experiencing problems because the frame has been formed with certain strut angles, making the attached shoulder bolt difficult to remove. It should be noted in the next frame study by considering the tolerance of each component in the frame external fixation to avoid backlash. The ring-strut angle should achieve the recommended angle. Stability of the frame to replace struts for longer or shorter struts.

# CONCLUSION

The use of standard components in OrthoLPPDUNS, spherical ball plain bearings GE 8E components can eliminate backlash and threaded rods ISO 724 to ensure uniform thread quality. The frame construction with ring-strut has a minimum safety factor value and the displacement value with IFS is below the maximum value. The range of the ring-strut angle in the frame still needs to be reexamined as it has not reached an angle of nearly 30 degrees. It is also important to consider changing the combination of strut sizes from medium to long or short with all configurations in the frame.

ACKNOWLEDGEMENT

The authors received no financial support for the research, and publication of this article. The authors thank Miss. Durkes for technical support and help in processing the data. The authors thank PT. SKF Indonesia Manufacturing through Mrs. Iffa Ardhiyana, in particular, for the donation of GE 8E and GE 10E Spherical Plain Bearing components to the joint system on the frame external fixation.

REFERENCES

Augat, P., Burger, J., Schorlemmer, S., Henke, T., Peraus, M., and Claes, L. (2003). “Shear movement at the fracture site delays healing in a diaphyseal fracture model.” Journal of Orthopaedic Research, Vol. 21, 1011-1017.

Behrens, F. (1989). “General theories and principles of external fixation.” Clinic Orthopaedics J., Vol. 241, 15-23.

Betts, S.C. and Muller, R. (2014). “Mechanical regulation of bone regeneration: theory, models, and experiments.” Frontiers in Endocrinology, Bone Research, 5, 211, 1-14. Accessed April 8, 2020, available from:https://www.ncbi.nlm.nih.gov/pmc/ articles/PMC4261821/pdf/fendo-05-00211.pdf

Birringer, R.P., Ganot, G.S., and James, B.A. (2016). “Failure Analysis of Internal Fixation Medical Devices: Overview and Case Studies.” Journal of Failure Analysis and Prevention, Vol. 16 (5), 849-857.

Bone Fixator. (2018). Taylor Spatial Frame. Accessed April 8, 2020, http://www.bonefixator.com/fracturefixation/taylor_ spatial_frame_1.htm l #figure2.

Dubowy, S.M. and and Nichols, L.R. (2018). Limb Lengthening Surgery: External Fixator. Accessed April 8, 2020, from: https://kidshealth.org/en/parents/limb-lengthen-fixator.html..

Drijber, F.L.I.P., Finlay, J.B., Moroz, T.K. and Rorabeck, C.H. (1990). “Sources of the slippage in the universal joints of the hoffman external fixator.” Med. & Biol. Eng. & Computer J., Vol. 28, 8-14.

Eralp, L., Bilen, F.E., Rozbruch, S.R., Kocaoglu, M., and Hammoudi, A. (2016). “External fixation reconstruction of the residual problems of benign bone tumours.” Strategies in Trauma and Limb Reconstruction, Vol. 11 (1), 37-49.

Evans, J. R. (2014). Quality & Performance Excellence. 7th ed. New York: South-Western College Pubisher.

Heidari, N., Hughes, A. and Atkins, R.M. (2013). “Intra-operative correction of taylor spatial frame without a computer.” J. Orthop Trauma. Vol. 27, e42–e44.

Henderson, R.C., Kemp, G.J., and Greene, W.B. (1992). “Adolescent tibia vara: alternatives for operative treatment.” J. Bone Joint Surg [Am]. Vol. 74A, 342-350.

Henderson, E.R., Feldman, D.S., Craig Lusk, C.L., van Bosse, H.J., Sala, D. and Kummer, F.J. (2008). “Conformational Instability of the Taylor Spatial Frame A Case Report and Biomechanical Study.” J. Pediatric Orthopeadic. Vol. 28, 471-477.

INA. (2008). INA Spherical Plain Bearings. Accessed April 8, 2020, from http://www.ina.de/content.ina.de/en/branches/ industry/solar_power_plants/products_solar/spherical_plain_bbearings/spherical_plain_bearings.jsp.

International Organization for Standardization. (1993). ISO general-purpose metric screw threads - Basic dimensions (ISO Standard No. 724).

Karidis, J.P. (2012). Orthopedic fixation device with zero backlash and adjustable compliance, and process for adjusting same. U.S. Patent 8,202,273 B2.

Keshet, D. and Eidelman, M. (2017). “Clinical utility of the taylor spatial frame for limb deformities.” Orthopedic Research and Reviews, Vol. 9, 51-61.

Nicholas (2010). Lean Production for Competitive Advantage: A Comprehensive Guide to Lean Methodologies and Management Practices, 1st ed. New York: CRC Press.

Norrish, A. (2013). Use of the Taylor Spatial Frame in Trauma. UK: Cambridge University Hospital.

Paley, D. (2002). Principles of Deformity Correction. Berlin Heidelberg: Springer-Verlag.

Paley, D. (2011). “History and science behind the six-axis correction external fixation devices in orthopaedic surgery.” Operative Techniques Orthopaedics, Vol. 21, 125-128.

Putra, I G.M. (2018). “Penyakit Blount - Diagnosis dan Tatalaksana. Fakultas Kedokteran Universitas Udayana, Denpasar, Bali, Indonesia.” Cermin Dunia Kedokteran 261. Vol. 45 (2), 112-114.

Seherr-Thoss, H.C., Schmelz, F. and Aucktor, E. (2006). Universal Joints and Driveshafts: Analysis, Design, Applications. German: Springer.

Sabharwal, S. (2016). Pediatric Lower Limb Deformities: Principles and Techniques of Management. New York: Springer International Publisher.

Shah, S.V., Saha, S.K., and Dutt, J.K. (2012). “Denavit-Hartenberg parameterization of euler angles.” Journal of Computational and Nonlinear Dynamics, Vol. 7 (2), 021006.

SKF. (2011). Catalogue: Spherical Plain Bearings and Rod Ends. Swedia: SKF Group.

Smith & Nephew. (2011). Taylor Spatial Frame External Fixator : Explore the Possibilities. USA: Smith & Nephew Inc.

Solomin, L.N. (2008). The Basic Principles of External Fixation Using the Ilizarov Device. Italia: Springer-Verlag Publisher.

Solomin, L.N. (2012). The Basic Principles of External Skeletal Fixation Using the Ilizarov and Other Devices. Italia: Springer-Verlag Publisher.

Solon, T. (2010). Backlash vs. Stiffness – What the different is and why you should care. Accessed April 8, 2020, available from: http://www.designworldonline.com/backlash-vs-stiffness-what-the-difference-is-and-why-you-should-care/

Steiner, M., Claes, L., Ignatius, A., Simon, U., and Wehner, T. (2014). “Disadvantages of interfragmentary shear on fracture healing-mechanical insights through numerical simulation.” Journal of Orthopaedic Research, Vol. 32 (7), 865-872.

Tan, B.B., Shanmugam, R., Gunalan, R., Chua, Y.P., Hossain, G., Saw, A. (2014). “A biomechanical comparison between taylor’s spatial frame and ilizarov external fixator.” Malaysia Orthopaedic J. Vol. 8 (2), 35-39.

Taylor, H.S. and J.C. (2000). Six Axis External Fixator Strut. US Patent 6030386.

Ullman, D.G. (1986). Mechanical Design Failure Analysis: with Failure Analysis System Software for the IBM Pc. New York: Marcel Dekker.

Vesali, F., Rezvani, M.A., and Kashfi, M. (2012). “Dynamics of universal joints, its failures and some propositions for practically improving its performance and life expectancy.” Journal of Mechanical Science and Technology. Vol. 26 (8), 2439-2449.

Wijaya, Y.I.D., Susmartini, S. and Herdiman, L. (2019). “Strut and ring connection system design on external fixator based on hexapod to increase bone flexibility reconstruction using finite element analysis.” IOP Conf. Series: Materials Science and Engineering, Vol. 495, 1-11.

# Footnotes

1	Industrial Engineering Department, Faculty of Engineering, Sebelas Maret University, Surakarta, 57126

2	Orthopedic Surgery and Traumatology Department, 
Dr. Soeradji Tirtonegoro Hospital, Klaten, 57424

a	email: lobesh@gmail.com

b	email: priadythama@gmail.com

c	email: susysus2011@gmail.com

d	email: priyandari@ft.uns.ac.id

e	email: idnsuci@gmail.com

f	email: durkesherlina@gmail.com

Diajukan: 21-04-2020	Diperbaiki: xx-xx-20xx

Disetujui: xx-xx-20xx

# Text boxes

Figure 4. The general operation of the frame external fixation.

Figure 6. The tracing twist strut becomes rough.

| a. Ilizarov external fixation | b. Monolateral external fixation | c. Taylor spatial frame external fixation |
| --- | --- | --- |
| Figure 1. Types of external fixation among orthopedic physicians |

Figure 2. Methodology of the redesign stage of external fixation.

Figure 3. Frame OrthoLPPDUNS with all custom component manufacturing.

Figure 5. The operations on frame external fixation.

Figure 7. The single universal joint standard.

Figure 7. The single universal joint standard.

Figure 8. The single custom universal joint.

Figure 9. The tracing on the strut screw wear out quickly.

Figure 10. Radial spherical plain bearing GE 8E (in mm).

Figure 11 The OrthoLPPDUNS design uses custom components.

Figure 12 The OrthoLPPDUNS design uses standard components.

Figure 13 The operation of the frame external fixation uses standard components.

Table 1. Comparison of characteristics and simulations on the stability parameters of the frame

| No | OrthoLPPDUNS | unit | Custom component | Standard component | Difference |
| --- | --- | --- | --- | --- | --- |
| 1 | Mass weight | kg | 0.855 | 0.749 | 0.106 |
| 2 | Width | mm | 217 | 225 | -8 |
| 3 | Height | mm | 174.7 | 139.74 | 34.96 |
| 4 | Safety factor | ul | 1.93 | 4.21 | -2.28 |
| 5 | Displacement | mm | 0.388 | 0.115 | 0.2727 |
| 6 | Maximum elongations ring and strut angle | degree | 0 | 42.08 |  |

Figure 14. Displacement on the external frame fixation with standard components.

Figure 15. The OrthoLPPDUNS frame uses several standard components.

| a. Custom universal joint components | b. Standard spherical plain bearing components |
| --- | --- |

Figure 16. Ring-strut angle on the joint system on the external fixation.